MEDICAL DEVICE SEMINAR  ·  PARTICIPANT WORKBOOK

Global Market Entry

United States  ·  European Union  ·  United Kingdom  ·  Japan

A technical overview of navigating MDR and UKCA requirements for Class II devices

How to use this workbook This is yours to write in. The note pages are keyed to the sections of the presentation so you can capture what matters to your product rather than transcribing slides. The two working sessions are done at your table during the seminar. The reference material at the back — the comparison matrix, the quick references, the glossary and the source list — is there for after you get home. The most useful page is the last one. Fill in three actions before you leave the room.

Contents

Agenda and learning objectives	2

Part 1 — Four markets, not four copies	3

Part 2 — EU MDR	4

Working session 1 — One device, four classifications	6

Part 3 — UKCA and the UK framework	7

Part 4 — United States and Japan	8

Working session 2 — Eighteen months of runway	9

Part 5 — Making it work	10

Reference: the four markets side by side	11

Reference: EU MDR at a glance	12

Reference: the UK position	13

Glossary	14

Primary sources — and what each is for	16

Your action plan	17

A note on currency Regulatory requirements in all four of these markets are moving. The UK framework is half-built, and the European Commission has a live proposal to amend the MDR. Dates and statuses in this workbook were checked against primary sources when it was prepared. Before relying on any of them — particularly for a submission or a board paper — confirm the current position at the source. The source list near the back tells you where to look.

AGENDA AND OBJECTIVES

Agenda and learning objectives

Run of show

| Time | Session |
| --- | --- |
| 0:00 | Welcome, objectives, room calibration |
| 0:04 | Part 1 — Four markets, not four copies |
| 0:15 | Part 2 — EU MDR |
| 0:37 | Working session 1 — One device, four classifications (10 min) |
| 0:47 | Part 3 — UKCA and the UK framework |
| 1:01 | Part 4 — United States and Japan |
| 1:12 | Working session 2 — Tabletop (10 min) |
| 1:22 | Part 5 — Making it work |
| 1:29 | Discussion and questions |

What you should be able to do afterwards

1.  Classify in four systems. Take one device and determine its regulatory class in the US, EU, UK and Japan — and explain why the answers differ rather than assuming one is wrong.

2.  Pick the right route. Choose a conformity assessment route under MDR and UK MDR 2002, and identify whether a Japanese device needs Ninsho or Shonin.

3.  Scope the technical file. Name what Annex II requires, what a 510(k) requires, and which parts of one package genuinely serve the other.

4.  Separate law from proposal. State what is in force, what is adopted but not yet applicable, and what is only proposed — and plan accordingly.

5.  Place the representatives. Explain what a US Agent, EU Authorised Representative, UK Responsible Person and Japanese MAH each carry.

6.  Sequence the programme. Argue for a market entry order with reasons, and identify the decisions that are expensive to reverse.

Three questions to hold in mind throughout

1. For my product, what is the class in each of the four markets — and can I show the reasoning? 2. Which of my decisions so far would be expensive to reverse? 3. What am I assuming travels between jurisdictions that may not?

PART ONE

Part 1 — Four markets, not four copies

“Class II” is a US term. It does not travel.

| Jurisdiction | Classification system | What a US Class II device becomes |
| --- | --- | --- |
| United States | Three classes, assigned by regulation per device type | Class II by definition — usually a 510(k) |
| European Union | Four classes, derived by applying the 22 rules of MDR Annex VIII to the intended purpose | Anywhere from I to III. The rules decide, not the US class |
| Great Britain | Four classes, derived from the MDD Annex IX rules as modified by Schedule 2A of the UK MDR 2002 | Usually the same as the EU — but not always, because GB uses the older rule set |
| Japan | Four classes tied to a JMDN code, with a parallel regulatory category | Determined by the JMDN code the product maps to |

Four regulatory models — four different questions

United States: FDA reviews you directly. The question is comparative: is this substantially equivalent to a predicate?

European Union: A notified body certifies you against the General Safety and Performance Requirements. The question is absolute: do you meet the requirements?

Great Britain: A UK Approved Body assesses against the UK MDR 2002 — the same model, an older rule set. And CE marking is still accepted.

Japan: A Registered Certification Body certifies where a certification standard exists; otherwise PMDA reviews and MHLW approves.

What transfers, and what does not

| Travels well — build once | Does not travel — budget for each |
| --- | --- |
| Bench and performance testing Biocompatibility (ISO 10993) The ISO 14971 risk management file Software lifecycle evidence (IEC 62304) One ISO 13485 quality system | The classification determination — four separate analyses The submission dossier itself Clinical evidence sufficiency Labelling and language The local representative |

Caveat: each regulator keeps its own list of recognised standards, and they are not synchronised. Check the list before you book the test.

The decision upstream of all four One sentence — the intended purpose — determines your class in four jurisdictions, your evidence burden and your route. Test: could two competent regulatory professionals, reading only that sentence, classify the device the same way in the EU? If not, it is not finished.

Notes — what applies to my product:

PART TWO

Part 2 — EU MDR

Annex VIII: what actually moves the class

Duration of contact: transient is under 60 minutes; short term is 60 minutes to 30 days; long term is over 30 days.

Invasiveness: body orifice versus surgically invasive, and whether it reaches the central circulatory or central nervous system.

Active or not: whether it depends on a source of energy, and whether it administers or exchanges energy, or supports or sustains life.

Clinical claim: diagnosis, monitoring, driving clinical management. This is where Rule 11 lives.

And apply the implementing rules in Chapter II first: accessories and software classify in their own right, and where several rules apply, the strictest wins.

Rule 11 — software

| What the software does | Class |
| --- | --- |
| Provides information used to take decisions for diagnostic or therapeutic purposes | IIa |
| … and those decisions may cause death or irreversible deterioration of health | III |
| … or may cause serious deterioration of health, or a surgical intervention | IIb |
| Intended to monitor physiological processes | IIa |
| … where the parameters are vital and variation could cause immediate danger | IIb |
| All other software | I |

Rule 11 turns on the CONSEQUENCE of the decision the software informs, not on how the software works. A simple rules engine and a deep learning model making the same clinical claim classify identically.

Class IIa and IIb — what differs

|  | Class IIa | Class IIb |
| --- | --- | --- |
| Technical documentation assessment | Representative basis — at least one device per category | Representative basis — at least one per generic device group, which is narrower |
| PSUR | At least every two years | At least annually |
| Implantables | Not usually applicable | Assessed device by device, with a listed set of exceptions |

The clinical evidence squeeze

Why the MDD equivalence route closed Under the MDD, a great many devices were carried by claiming equivalence to a competitor's device and citing its published literature. MDR did not ban equivalence — it made it impractical. Article 61(5) requires, where the equivalent device belongs to another manufacturer, a contract giving you full and ongoing access to their technical documentation. Your competitor will not sign it. The practical consequence: many devices that needed no clinical study under the MDD need PMCF data, a registry or an actual investigation under MDR. This is the largest gap between an FDA clearance and a CE certificate.

The transition — Regulation (EU) 2023/607

| Deadline | Applies to |
| --- | --- |
| 31 Dec 2027 | Class III, and Class IIb implantables — except sutures, staples, dental fillings, dental braces, tooth crowns, screws, wedges, plates, wires, pins, clips and connectors |
| 31 Dec 2028 | Other Class IIb, all Class IIa, and Class I placed on the market sterile, with a measuring function, or reusable surgical instruments |

Conditions, all of which must hold: the device is a legacy device covered by a certificate or declaration of conformity issued under the directives before 26 May 2021; it still complies with that directive; there is no significant change in design or intended purpose; it presents no unacceptable risk; and by 26 May 2024 the manufacturer had an MDR Article 10(9) quality system and had lodged a formal application with a notified body, with a signed agreement by 26 September 2024.

The extension is not a delay you can take. It is conditional relief you either qualified for by those 2024 dates, or did not.

And what is only proposed

COM(2025) 1023 final, 16 December 2025, procedure 2025/0404(COD) — a Commission proposal to simplify MDR and IVDR. Among other things it would reduce notified body involvement for Class IIa and IIb, remove the five-year certificate validity limit, loosen the clinical evidence and equivalence rules, adapt some classification rules downward including for software, and extend non-critical vigilance reporting from 15 to 30 days.

It is a proposal under the ordinary legislative procedure. It is not law, it may change substantially, and it may fail. Track it; plan on current law until it is published in the Official Journal.

WORKING SESSION 1  ·  10 MINUTES

One device, four classifications

The device A wearable continuous vital-signs monitor for hospital ward patients. · An adhesive chest patch measures heart rate, respiratory rate and skin temperature continuously. · It transmits to a bedside gateway and then to a server on the hospital network. · Software on that server runs an early-warning algorithm and raises an alert to the nursing station when it detects signs of clinical deterioration. · The patch is worn for up to 5 days and is single-patient use. · It is cleared in the US as a Class II device.

1.  Split the product into its regulated parts

How many devices are there here, and is the algorithm one of them?

| Part | Is it a device in its own right? | Why |
| --- | --- | --- |

2.  Classify under MDR Annex VIII

| Part | Rule(s) applied | Class |
| --- | --- | --- |

3.  Now classify for Great Britain — remember GB uses the MDD rules

Do you get the same answer? If not, why not?

4.  Which EU conformity assessment route, and who assesses what?

5.  What would you need in Japan — and what must you know first?

PART THREE

Part 3 — UKCA and the UK framework

The United Kingdom is not one market

Great Britain: England, Scotland and Wales. Governed by the UK Medical Devices Regulations 2002 as amended. UKCA applies here — and CE marking is also currently accepted.

Northern Ireland: under the Windsor Framework, EU rules continue to apply. A device needs EU MDR conformity and CE marking. UKCA alone is not valid there.

The whole UK: MHRA registration applies UK-wide whatever the marking, and a non-UK manufacturer needs a UK Responsible Person.

The consequence: a CE marked device can be sold across the whole United Kingdom today. A UKCA marked device cannot — it does not satisfy Northern Ireland.

The dates that govern Great Britain

| Your device carries | Accepted in Great Britain until |
| --- | --- |
| CE marking under EU MDR | 30 June 2030 |
| CE marking under the MDD or AIMDD | the sooner of certificate expiry or 30 June 2028 |
| CE marking under the IVDD | the sooner of certificate expiry or 30 June 2030 |
| UKCA marking | no end date — and not mandatory |

On a legacy MDD certificate there are two clocks, and the earlier one governs. The government has also said it will consult on recognising CE marked devices indefinitely.

What UKCA actually requires

Not a British version of MDR. The UK MDR 2002 implemented the old European directives, and Schedule 2A explains how the directive annexes are read in a UK context. Classification is under MDD Annex IX, not MDR Annex VIII. An MDR technical file already exceeds what UKCA asks for in most respects — the file is not the obstacle, the assessment is.

Where the EU and UK have genuinely diverged

| Topic | EU MDR | Great Britain |
| --- | --- | --- |
| Classification | MDR Annex VIII — 22 revised rules | MDD Annex IX via Schedule 2A |
| Software | Rule 11 classifies it in its own right, commonly IIa or IIb | No equivalent rule — classified under older provisions and can land lower |
| Clinical evidence | Article 61 and Annex XIV, tightened equivalence | Directive-era expectations — materially lighter |
| Named role | Article 15 PRRC | No direct equivalent; the UK RP is a different role |
| Database | EUDAMED | MHRA registration — separate system |

The post-market surveillance regulations

In force since 16 June 2025 The Medical Devices (Post-market Surveillance Requirements) Regulations 2024 are the first substantial piece of the new UK framework to become law. They introduce a statutory PMS system, defined reporting to the MHRA, and periodic reporting. They attach to placing a device on the UK market, not to the conformity route — so a CE marked device sold in Great Britain is in scope. This breaks the assumption that relying on CE marking means no UK regulatory work. Registration and a Responsible Person can be delegated; an operating PMS system cannot.

PART FOUR

Part 4 — United States and Japan

The 510(k) in four steps

1.  Find a predicate — a legally marketed device to which you claim substantial equivalence. Check its recall history before you build around it.

2.  Establish substantial equivalence: same intended use, and either the same technological characteristics or different ones that do not raise different questions of safety and effectiveness.

3.  Assemble the submission — device description, predicate comparison, performance testing, biocompatibility, sterilisation, software, labelling — via eSTAR. Then survive acceptance review and substantive review, remembering the clock pauses whenever FDA asks a question.

And check whether you need one at all: many Class II device types are exempt from premarket notification, subject to limitations set out in the classification regulation. Read the whole regulation — the special controls are binding.

Japan: the route depends on your code

| Class | Category | Premarket route |
| --- | --- | --- |
| I | General medical device | Todokede — notification to the PMDA |
| II | Controlled medical device | WITH a certification standard: Ninsho, by a Registered Certification Body. WITHOUT one: Shonin, approval of the MHLW on PMDA review |
| III | Specially controlled | The same split — Ninsho where a certification standard exists, otherwise Shonin |
| IV | Specially controlled | Shonin — approval of the MHLW |

Both determinants are lookups, not arguments: the class follows the JMDN code, and the route follows whether a certification standard exists for it. Until you have done that research you cannot forecast Japan at all.

Who holds what in Japan

Marketing Authorization Holder: a company established in Japan holds the authorisation and carries quality and safety responsibility. A foreign manufacturer cannot hold it directly.

Designated MAH: holds it on your behalf where you have no Japanese entity. Often a distributor — and the choice has long consequences.

Foreign Manufacturer Registration: the overseas site is registered separately, and that workstream runs in parallel with the product application.

The commercial trap If your distributor is your DMAH, your authorisation sits in their name. Changing partner then means transferring or re-obtaining the approval, and they have no incentive to be quick. Negotiate transfer terms into the original agreement: nobody objects at the start of a relationship; everybody objects at the end.

One quality system, four audiences

| Jurisdiction | Requirement | MDSAP accepted? |
| --- | --- | --- |
| United States | QMSR — 21 CFR 820 on ISO 13485:2016 | Yes, for routine inspections |
| European Union | MDR Article 10(9); ISO 13485 in practice | No — your notified body audits you |
| Great Britain | Per the UK MDR route; ISO 13485 | Not as a substitute — check policy |
| Japan | QMS Ordinance — ISO 13485 plus additions | Participates — confirm scope |

Build to ISO 13485 as the base and layer the jurisdiction-specific additions on top. Running parallel quality systems is the most expensive avoidable mistake in this field.

WORKING SESSION 2  ·  10 MINUTES  ·  TABLETOP

Eighteen months of runway

The situation You are the regulatory lead. Your company has the wearable monitor from the first session. It is cleared in the US and selling modestly. You have 18 months of runway, one regulatory hire you have not yet made, and a board that wants international revenue in this financial year. Your CEO has decided you will “do Europe, the UK and Japan”. Your clinical evidence is bench testing and a 60-patient usability study. You have no notified body contract. Your Japanese distributor has offered to act as DMAH at no cost.

Round 1 — sequence it

Which market first, and why?  What do you tell the board about revenue this year?  What does your one hire need to be good at?

Round 2 — find what breaks

Where does the clinical evidence fall short, and for which market?  What is the earliest realistic EU date?  What would you do about the UK specifically?

Round 3 — the decisions you cannot reverse

What do you say to the Japanese distributor?  What must be decided in the next 30 days?  What would you refuse to do, even under board pressure?

PART FIVE

Part 5 — Making it work

One technical file, four submissions

1.  Hold evidence separately from submissions. Test reports, the risk file, clinical data and design records are durable assets. A 510(k), an Annex II file, a UKCA file and a Ninsho application are views of those assets assembled for different readers.

2.  Index against the strictest requirement set. Build the index around MDR Annex II and the GSPR checklist. Everything else is then a subset or a reformatting, not new work.

3.  Keep four classification rationales. One document per jurisdiction, each citing its own rules, each dated. Two hours of work that answers the first question every assessor asks.

4.  Track standards and editions in one register. Which standard, which edition, which report, which regulator recognises it. This is where reuse either works or quietly fails.

Which market first?

| The case for Europe first | The case against — and what to do instead |
| --- | --- |
| It is the most demanding, so everything else becomes a subset The MDR file over-satisfies UKCA and much of it serves Japan A CE certificate currently reaches the whole UK Notified body queues are long, so starting early has option value Take this route when you have the clinical evidence, or the runway to generate it | It is the slowest and most expensive first step The clinical evidence gap lands on the critical path Revenue arrives last, which is the opposite of what most boards need Instead: sequence for what unblocks revenue soonest while building toward the strictest file — and start the European evidence in parallel from day one |

Whichever order you choose: start the clinical evidence strategy on day one, in every scenario. Almost everything else can be accelerated with money. Clinical evidence cannot.

Five things worth remembering

1.  “Class II” is a US term and it does not travel. Four jurisdictions, four classification systems, four determinations from the same intended purpose.

2.  The four regulators ask different questions. Substantial equivalence, conformity with the GSPRs, conformity with the directives, conformity with a standard.

3.  Clinical evidence is the incompressible item. The largest gap between an FDA clearance and a CE certificate, and the only workstream money cannot accelerate.

4.  Know what is in force, adopted, and merely proposed. The EU simplification package is a proposal. The UK pre-market framework is an intention. The UK post-market regulations are law.

5.  The expensive decisions are made early and cheaply. The intended purpose sentence, the predicate, the notified body, the DMAH agreement.

REFERENCE

The four markets side by side

|  | United States | European Union | Great Britain | Japan |
| --- | --- | --- | --- | --- |
| Class basis | By regulation, per device type | MDR Annex VIII rules | MDD Annex IX rules | Follows the JMDN code |
| Who assesses | FDA directly | A notified body | A UK Approved Body | An RCB, or PMDA/MHLW |
| Premarket route | 510(k), De Novo or PMA — or exempt | Annex IX, X or XI | UKCA — or rely on CE for now | Todokede, Ninsho or Shonin |
| Core question | Substantially equivalent to a predicate? | Does it meet the GSPRs? | Does it meet the directive requirements? | Does it meet the standard for its code? |
| Clinical evidence | Often none for a 510(k) | Clinical evaluation for every device | Directive-era — lighter | Foreign data accepted; bridging may be needed |
| Local role | US Agent | Authorised Rep — jointly liable | UK Responsible Person | MAH or DMAH — holds the authorisation |
| Registration | Establishment registration and listing, annual | EUDAMED | MHRA — whatever the marking | Via the MAH |
| QMS | QMSR — 820 on ISO 13485 | Article 10(9); ISO 13485 | Per the route; ISO 13485 | QMS Ordinance — 13485 plus |
| MDSAP accepted | Yes | No | Not as a substitute | Participates |
| Post-market | MDR reporting; 806 corrections | Annex III: PMS, PMCF, PSUR | PMS regs since Jun 2025 | Via the MAH, on a QMS cycle |
| The trap | Assuming clearance means the file travels | Finding the clinical gap after contracting an NB | Building UKCA you may never need | Letting your distributor hold the authorisation |

Read the “core question” row across Four different questions. That row alone explains why the evidence packages differ, why an FDA clearance does not carry Europe, and why the same device can be straightforward in one market and a programme in another. And notice the bottom row: each market has a characteristic expensive mistake, and three of the four are decisions made early, cheaply, and usually without anyone realising a decision is being made.

REFERENCE

EU MDR at a glance

The Annex II technical file

| Element | What makes it adequate |
| --- | --- |
| Device description and specification | Variants, accessories, Basic UDI-DI, intended purpose, intended population and user, principles of operation. |
| Information supplied by the manufacturer | Labels and IFU in the languages of every member state where the device is made available. |
| Design and manufacturing information | Design stages, manufacturing processes, and the sites — including suppliers and subcontractors. |
| GSPR checklist | Every Annex I requirement, applicability, method of demonstration, and a precise cross-reference to evidence. |
| Benefit-risk and risk management | The ISO 14971 file, with controls traced to verification that they are implemented and effective. |
| Verification and validation | Bench, biocompatibility, electrical safety, software, usability, stability and shelf life, and the clinical evaluation report. |
| Annex III — post-market surveillance | A separate annex and a separate obligation: the PMS plan, and the PSUR or PMS report by class. |

Build the GSPR checklist first, not last

For every Annex I requirement: does it apply, by what method will you demonstrate conformity, and exactly which document and section holds the evidence. Built early it is a gap analysis that tells you what to test. Built at the end it is a discovery exercise that tells you what you forgot.

Accountability and traceability

Article 15 — PRRC: a named person responsible for regulatory compliance. Micro and small enterprises need not employ one but must have one permanently and continuously at their disposal.

Article 11 — Authorised Representative: required for any manufacturer outside the EU, acting under a written mandate. Article 11(5) makes them jointly and severally liable with you for defective devices.

Basic UDI-DI: identifies the model, appears on the declaration of conformity and in EUDAMED, and is NOT on the label. The UDI-DI plus UDI-PI is the carrier on the device.

Post-market obligations

PMS plan: documented, proactive and systematic; assessed by the notified body before you sell anything.

PMCF: the mechanism that keeps the clinical evaluation current. Not doing it requires a justification that gets read carefully.

PSUR: at least every two years for Class IIa; at least annually for Class IIb and III.

Vigilance: serious incidents and field safety corrective actions within defined timescales, plus trend reporting.

The connection people miss PMCF data feeds the clinical evaluation, which supports conformity, which is what the certificate attests. A thin PMCF programme is not only an audit risk — it erodes the evidence your certificate rests on, and you discover that at renewal when there is no time.

REFERENCE

The UK position

Decision aid: do you need UKCA?

| Question | If yes | If no |
| --- | --- | --- |
| Do you hold, or expect, a CE certificate under EU MDR? | It reaches Great Britain until 30 June 2030, and Northern Ireland indefinitely. UKCA adds nothing you do not already have. | UKCA may be your route to Great Britain — check Approved Body capacity for your device code first. |
| Are you on a legacy MDD certificate? | Your GB access ends at the sooner of certificate expiry or 30 June 2028. Two clocks; the earlier governs. | Then the MDR dates above apply. |
| Do you sell in Northern Ireland? | You need CE marking regardless. UKCA alone does not satisfy Northern Ireland. | Great Britain alone is a smaller market than the CE route already reaches. |
| Is your business UK-only? | UKCA may well be the right and cheaper answer. | Weigh UKCA against simply waiting — the government has said it will consult on indefinite CE recognition. |

What applies whatever marking you carry

A UK Responsible Person: required if you are established outside the UK. A UK address, a written agreement, and real standing access to your technical documentation.

MHRA registration: applies across the United Kingdom regardless of marking, and is made by the UK RP where the manufacturer is outside the UK.

Post-market surveillance: the Medical Devices (Post-market Surveillance Requirements) Regulations 2024, in force since 16 June 2025. These attach to placing the device on the market, not to the conformity route.

Vigilance reporting: serious incidents and field safety corrective actions to the MHRA.

The state of the framework

| What | Status |
| --- | --- |
| Post-market surveillance regulations | IN FORCE since 16 June 2025 |
| A strengthened pre-market regime | Signalled, not law. Timing has moved before — treat any date you hear as provisional |
| Indefinite recognition of CE marking | Government has said it will consult. Not decided |
| International reliance routes | The MHRA has been building these. Watch — it is the likely shape of the future |

The correct response to a half-built framework Comply fully with what is in force, build nothing speculative for what is not, and make sure someone owns watching it — by name, not by committee. The manufacturers who lose money here are the ones who either ignore the UK entirely or try to build for a regime that has not been written.

REFERENCE

Glossary

| Term | Meaning |
| --- | --- |
| 510(k) | US premarket notification demonstrating substantial equivalence to a legally marketed predicate device. |
| AIMDD | The Active Implantable Medical Devices Directive, 90/385/EEC — one of the directives carried into UK law. |
| Annex I (MDR) | The General Safety and Performance Requirements: 23 requirements in three chapters, and the spine of the technical file. |
| Annex II (MDR) | The contents of the technical documentation. |
| Annex III (MDR) | The technical documentation on post-market surveillance — a separate annex and a separate obligation. |
| Annex VIII (MDR) | The classification rules: 22 rules, applied to the intended purpose. |
| Annex IX / X / XI (MDR) | The conformity assessment routes: QMS plus technical documentation assessment; type-examination; production conformity. |
| Approved Body | A conformity assessment body designated by the MHRA to assess against the UK MDR 2002. The UK counterpart to a notified body. |
| Authorised Representative | The EU role under MDR Article 11, required for manufacturers outside the EU. Jointly and severally liable for defective devices under Article 11(5). |
| Basic UDI-DI | The identifier for a group of devices sharing intended purpose, class and design. Appears on the declaration of conformity and in EUDAMED — not on the label. |
| CE marking | The mark indicating conformity with the applicable EU legislation. |
| Clinical evaluation | The systematic assessment of clinical data required under MDR Article 61 and Annex XIV for every device. |
| De Novo | The US route for a novel low-to-moderate risk device with no predicate, which creates a new classification. |
| DMAH | Designated Marketing Authorization Holder — a Japanese entity holding the marketing authorisation on a foreign manufacturer's behalf. |
| eSTAR | FDA's electronic submission template, mandatory for 510(k) submissions. |
| EUDAMED | The European database for medical devices: actors, devices, certificates, clinical investigations, vigilance and market surveillance. |
| GSPR | General Safety and Performance Requirements — MDR Annex I. |
| JMDN | Japanese Medical Device Nomenclature. Your code determines your class and therefore your route. |
| MAH | Marketing Authorization Holder — the Japanese entity that holds the authorisation and carries quality and safety responsibility. |
| MDCG | Medical Device Coordination Group — publishes the guidance that interprets MDR in practice. |
| MDD | The Medical Devices Directive, 93/42/EEC. Superseded in the EU by MDR; still the substantive basis of UK requirements. |
| MDR | Regulation (EU) 2017/745 on medical devices. |
| MDSAP | Medical Device Single Audit Program. One audit accepted by several regulators — but not by the EU. |
| MHRA | The UK Medicines and Healthcare products Regulatory Agency. |
| NANDO | The EU database of designated notified bodies and their scopes. Check it before contracting. |
| Ninsho | Certification by a Japanese Registered Certification Body, for Class II and III devices where a certification standard exists. |
| Notified Body | A private organisation designated by an EU member state to carry out conformity assessment under MDR. |
| PMCF | Post-market clinical follow-up — the mechanism that keeps the clinical evaluation current. |
| PMD Act | Japan's Pharmaceutical and Medical Device Act. |
| PRRC | Person Responsible for Regulatory Compliance, under MDR Article 15. |
| PSUR | Periodic Safety Update Report. Every two years for Class IIa; annually for IIb and III. |
| QMSR | FDA's Quality Management System Regulation — 21 CFR 820 incorporating ISO 13485:2016, in force since 2 February 2026. |
| RCB | Registered Certification Body — the Japanese third-party certifier for the Ninsho route. |
| Rule 11 | The MDR software classification rule. It turns on the consequence of the decision the software informs. Great Britain has no equivalent. |
| Schedule 2A | The schedule inserted into the UK MDR 2002 setting out how the directive annexes are read in a UK context. |
| Shonin | Approval by the MHLW on PMDA review — the Japanese route for Class IV, and for Class II and III where no certification standard exists. |
| Substantial equivalence | The comparative standard a 510(k) must meet against its predicate. |
| Todokede | Notification to the PMDA — the Japanese route for Class I devices. |
| UDI | Unique Device Identification. UDI-DI plus UDI-PI form the carrier on the device and its packaging. |
| UK MDR 2002 | The Medical Devices Regulations 2002, as amended — the law governing Great Britain. |
| UK Responsible Person | The UK role required for manufacturers established outside the United Kingdom. Separate from an EU Authorised Representative. |
| UKCA | UK Conformity Assessed marking. Available, and currently not mandatory. |
| Windsor Framework | The arrangement under which EU rules continue to apply to goods in Northern Ireland. |

REFERENCE

Primary sources — and what each is for

Primary sources only. Dates and statuses were checked against these when this workbook was prepared. Several are moving — confirm before relying on any of them.

European Union

| Source | Read it for |
| --- | --- |
| Regulation (EU) 2017/745 (MDR), consolidated text on EUR-Lex | The actual articles and annexes. More readable than people expect, and summaries lose the conditions. |
| Regulation (EU) 2023/607 | The extended transition deadlines and — more importantly — the conditions attached to them. |
| COM(2025) 1023 final, procedure 2025/0404(COD) | The simplification proposal. Check its current stage before citing it — it is not law. |
| MDCG guidance, especially 2020-5 and 2020-6 | How equivalence is assessed, and what legacy clinical data can support. |
| The Commission's EUDAMED pages | Which modules are mandatory today. This has changed more than once. |
| NANDO | Notified body designations and scopes. Check a body covers your device code before contracting. |

United Kingdom

| Source | Read it for |
| --- | --- |
| gov.uk — Regulating medical devices in the UK | The authoritative statement of CE acceptance dates, registration and the UK Responsible Person. |
| The Medical Devices Regulations 2002 on legislation.gov.uk | The law itself, including Schedule 2A and how the directive annexes are read. |
| The Medical Devices (Post-market Surveillance Requirements) Regulations 2024 | What the PMS obligations in force since 16 June 2025 actually require. |
| The MHRA's list of UK Approved Bodies | Who is designated, and for what scope. |

United States and Japan

| Source | Read it for |
| --- | --- |
| FDA classification database and the product code | Your class, the applicable regulation, and the special controls — which are binding, not advisory. |
| FDA guidance on the 510(k) programme, De Novo and the QMSR | The route, the evidence expectations, and the quality system requirement in force since February 2026. |
| PMDA — Regulations and Approval/Certification of Medical Devices | The class-to-route map, in the regulator's own words. Japan explains its system better than any third party. |
| MDSAP programme documents | Which regulators accept what, and precisely what it does and does not replace. |

THE PAGE THAT MATTERS

Your action plan

Fill in at least three rows before you leave the room. A commitment made here is the only part of a seminar that reliably survives the journey back to the office.

| What I will do | Who owns it | By when |
| --- | --- | --- |

Ten questions to ask your team on Monday

1.  What is our device's class in each of the four markets, and where is the reasoning written down?

2.  Does our intended purpose statement say what we actually claim — and does it match our website?

3.  If we are on a legacy MDD certificate: when does it expire, and what does that do to Great Britain?

4.  Did we pass the 2024 gates for the extended transition, and can we evidence it?

5.  What is our clinical evidence strategy for MDR, and when did it start?

6.  Have we read our EU Authorised Representative mandate, including the liability clause?

7.  Do we have a UK Responsible Person, and do they hold our current technical documentation?

8.  Are we running one quality system or several — and if several, why?

9.  Who holds our marketing authorisation in Japan, and what does the agreement say about transferring it?

10.  Which regulator's recognised-standard list did we check before booking our last test?

If you do only one thing Write down your classification reasoning for each of the four markets, with the rules cited and the date. It takes a couple of hours, it answers the first question every assessor asks, and it forces the conversation about intended purpose that everything else depends on.