QUICK REFERENCE  ·  SIDE 1  ·  THE FOUR MARKETS

Classify, then route

“Class II” is a US term. It does not travel. Four jurisdictions, four determinations from the same intended purpose — write the reasoning down for each.

The four systems

|  | United States | European Union | Great Britain | Japan |
| --- | --- | --- | --- | --- |
| Class basis | Assigned by regulation, per device type (21 CFR 862–892) | Derived from the 22 rules of MDR Annex VIII | Derived from MDD Annex IX via Schedule 2A, UK MDR 2002 | Follows the JMDN code |
| Who assesses | FDA directly | A notified body | A UK Approved Body | A Registered Certification Body, or PMDA/MHLW |
| Route | 510(k), De Novo or PMA — or exempt | Annex IX, X or XI | UKCA — or rely on CE marking for now | Todokede, Ninsho or Shonin |
| Core question | Substantially equivalent to a predicate? | Does it meet the GSPRs? | Does it meet the directive requirements? | Does it meet the standard for its code? |
| Clinical evidence | Often none for a 510(k) | Clinical evaluation for EVERY device | Directive-era — lighter | Foreign data accepted; bridging may be needed |
| Local role | US Agent — largely a mailbox | Authorised Rep — jointly and severally liable, Art 11(5) | UK Responsible Person — separate appointment | MAH or DMAH — holds the authorisation |
| Registration | Establishment registration and listing, annual | EUDAMED | MHRA — whatever the marking | Via the MAH |
| QMS | QMSR — 21 CFR 820 on ISO 13485 | Art 10(9); ISO 13485 in practice | Per the route; ISO 13485 | QMS Ordinance — 13485 plus additions |
| MDSAP | Accepted for routine inspections | NOT accepted | Not as a substitute | Participates — confirm scope |
| The trap | Assuming clearance means the file travels | Finding the clinical gap after contracting an NB | Building UKCA you may never need | Letting your distributor hold the authorisation |

Japan: the route depends on your code

| I | General | Todokede — notification to the PMDA |
| --- | --- | --- |
| II | Controlled | WITH a certification standard: Ninsho, by a Registered Certification Body. WITHOUT one: Shonin, approval of the MHLW on PMDA review |
| III | Specially controlled | The same split — Ninsho where a certification standard exists, otherwise Shonin |
| IV | Specially controlled | Shonin — approval of the MHLW |

Both determinants are lookups, not arguments. Research the JMDN code first — until you have, you cannot forecast Japan at all.

The 510(k) in four steps

1  Find a predicate — check its recall history.   2  Establish substantial equivalence — same intended use, and either the same technological characteristics or different ones not raising different questions of safety and effectiveness.   3  Assemble and submit via eSTAR.   4  Acceptance review, then substantive review — the clock pauses whenever FDA asks a question.

And check whether you need one at all: many Class II device types are exempt, subject to the limitations in the classification regulation. The special controls in that regulation are binding.

What transfers between jurisdictions

| Travels well | Bench and performance testing · biocompatibility (ISO 10993) · the ISO 14971 risk file · software lifecycle evidence (IEC 62304) · one ISO 13485 quality system |
| --- | --- |
| Does not | The classification determination · the submission dossier itself · clinical evidence sufficiency · labelling and language · the local representative |

Caveat: each regulator keeps its own list of recognised standards, and the editions are not synchronised. Check the list before you book the test.

QUICK REFERENCE  ·  SIDE 2  ·  EU MDR AND THE UK

Routes, dates and divergence

MDR Rule 11 — software

| Provides information used to take decisions for diagnostic or therapeutic purposes | IIa |
| --- | --- |
| … and those decisions may cause death or irreversible deterioration of health | III |
| … or may cause serious deterioration of health, or a surgical intervention | IIb |
| Intended to monitor physiological processes | IIa |
| … where the parameters are vital and variation could cause immediate danger | IIb |
| All other software | I |

Rule 11 turns on the CONSEQUENCE of the decision the software informs, not on how the software works. Great Britain has no equivalent rule — the same software can classify lower there.

Annex VIII — what moves the class

Duration: transient under 60 min · short term 60 min to 30 days · long term over 30 days.   Invasiveness: orifice versus surgically invasive, and whether it reaches the central circulatory or central nervous system.   Active: whether it depends on energy, and whether it administers energy or sustains life.   Clinical claim: diagnosis, monitoring, driving management.

Apply the Chapter II implementing rules first: accessories and software classify in their own right, and where several rules apply the strictest wins.

The MDR transition — Regulation (EU) 2023/607  [IN FORCE]

| 31 Dec 2027 | Class III, and Class IIb IMPLANTABLES — except sutures, staples, dental fillings, dental braces, tooth crowns, screws, wedges, plates, wires, pins, clips and connectors |
| --- | --- |
| 31 Dec 2028 | Other Class IIb, all Class IIa, and Class I placed on the market sterile, with a measuring function, or reusable surgical instruments |

Conditions: a legacy device certified under the directives before 26 May 2021 · still compliant with that directive · no significant change in design or intended purpose · no unacceptable risk · an MDR Art 10(9) QMS and a formal notified body application by 26 May 2024, with a signed agreement by 26 Sept 2024. The extension is conditional relief you either qualified for, or did not.

Great Britain — the dates

| CE marking under EU MDR | Accepted until 30 June 2030 |
| --- | --- |
| CE marking under MDD or AIMDD | Accepted until the SOONER of certificate expiry or 30 June 2028 — two clocks, the earlier governs |
| CE marking under the IVDD | Accepted until the sooner of certificate expiry or 30 June 2030 |
| UKCA marking | No end date — and not mandatory. The MHRA has consulted on indefinite recognition of CE marking; the response is pending |

Do you need UKCA?

A CE marked device reaches the WHOLE United Kingdom today, because EU rules still apply in Northern Ireland under the Windsor Framework. A UKCA marked device reaches Great Britain only. So for a manufacturer with an EU route, UKCA is currently additional work to reach a smaller market.

Consider UKCA anyway if: you are a UK-only business; you cannot get a notified body slot but can get an Approved Body one; or you want insurance against the recognition policy changing. Whatever you carry, MHRA registration and a UK Responsible Person apply — and so do the Post-market Surveillance Regulations 2024, in force since 16 June 2025.

EU and UK divergence

| Topic | EU MDR | Great Britain |
| --- | --- | --- |
| Classification | MDR Annex VIII — 22 revised rules | MDD Annex IX via Schedule 2A |
| Software | Rule 11, in its own right | No equivalent rule — can land lower |
| Clinical evidence | Art 61, Annex XIV; tightened equivalence | Directive-era — materially lighter |
| Named role | Art 15 PRRC | No direct equivalent |
| Database | EUDAMED | MHRA registration — separate system |

Proposed, and NOT law

COM(2025) 1023 final, 16 December 2025, procedure 2025/0404(COD) — a Commission proposal to simplify MDR and IVDR. It would reduce notified body involvement for Class IIa and IIb, remove the five-year certificate validity limit, loosen clinical evidence and equivalence rules, adapt some classification rules downward including for software, relax the PRRC requirements, and extend non-critical vigilance reporting from 15 to 30 days. Track it; plan on current law until it is published in the Official Journal.