MEDICAL DEVICE SEMINAR  ·  PARTICIPANT WORKBOOK

Software Validation & Cybersecurity

Addressing the challenges of connected devices and Software as a Medical Device

How to use this workbook This is yours to write in. The note pages are keyed to the sections of the presentation so you can capture what matters to your product rather than transcribing slides. The two working sessions are done at your table during the seminar. The reference material at the back — the checklists, glossary and standards list — is there for after you get home. The single most useful page is the last one. Fill in three actions before you leave the room.

Contents

Agenda and learning objectives	2

Part 1 — The new operating environment	3

Part 2 — Software validation	4

Working session 1 — Classify and scope	5

Part 3 — Cybersecurity	6

Working session 2 — Monday morning, 07:40	7

Part 4 — Making it work	9

Readiness self-assessment	10

Reference: documentation level and safety class	11

Reference: common findings that recur	12

Reference: security baseline checklist	13

Glossary	15

Standards and guidance — and what each is for	17

Your action plan	18

A note on currency Regulatory guidance in this field is revised frequently. Dates given in this workbook are stated as at the time of preparation. Before relying on any document cited here — particularly for a submission — confirm the current version on the issuing body's website.

AGENDA AND OBJECTIVES

Agenda and learning objectives

Run of show

| Time | Session |
| --- | --- |
| 0:00 | Welcome, objectives, room calibration |
| 0:09 | Part 1 — The new operating environment |
| 0:31 | Part 2 — Software validation |
| 1:21 | Working session 1 — Classify and scope (12 min) |
| 1:33 | Part 3 — Cybersecurity |
| 2:17 | Working session 2 — Tabletop (15 min) |
| 2:32 | Part 4 — Making it work |
| 2:45 | Discussion and questions |

What you should be able to do afterwards

1.  Classify correctly. Decide whether a software function is SaMD, software in a device, or outside the device definition — and assign an IEC 62304 safety class and an FDA documentation level you can defend.

2.  Build the evidence. Name the validation deliverables a submission and an inspection actually require, and explain what makes each adequate rather than merely present.

3.  Run security as engineering. Describe the Secure Product Development Framework, run a threat model, and distinguish security risk from safety risk without collapsing one into the other.

4.  Meet section 524B. State the four statutory cybersecurity obligations for cyber devices and what evidence demonstrates each.

5.  Handle a vulnerability. Work a postmarket disclosure through triage, controlled-versus-uncontrolled risk, reporting decisions and customer communication.

6.  Find your gaps. Leave with a prioritised list of what your own programme is missing and a sequence to close it.

Three questions to hold in mind throughout

1. Which of these obligations already applies to a product we have shipped? 2. Who in our organisation owns this — by name, not by function? 3. What would we do if this happened on a Friday afternoon?

PART ONE

Part 1 — The new operating environment

The convergence

Three forces are now true at once, for the same product, across a service life measured in years: the device became connected; the software became the device; and the adversary acquired economic motive. None is new alone. The combination is.

What changed recently

| When | What | Why it matters |
| --- | --- | --- |
| Mar 2023 | Section 524B added to the FD&C Act | Cybersecurity became statutory, not guidance. FDA may refuse to accept a deficient submission. |
| Jun 2023 | Premarket software guidance finalised | Level of Concern retired; Basic and Enhanced documentation levels introduced. |
| Jun 2025 | Premarket cybersecurity guidance reissued, superseding the Sept 2023 edition | SPDF, architecture views, SBOM expectations, testing evidence, plus a section addressing s.524B directly. |
| Dec 2024 | PCCP guidance finalised for AI-enabled device software functions | Pre-authorised change control for software that is expected to change. |
| Feb 2026 | QMSR in force; QSIT retired | 21 CFR 820 incorporates ISO 13485:2016. Substance carries over; your citations change. |

Three kinds of risk — three different questions

|  | Safety risk | Security risk | Quality / process risk |
| --- | --- | --- | --- |
| Standard | ISO 14971 | AAMI SW96, IEC 81001-5-1 | ISO 13485 / QMSR |
| Asks | Could this device harm a patient? | Could someone make it harm a patient, or deny its use? | Could our process let a defect through? |
| Likelihood | Probability — estimable | Exploitability — not a frequency | Process capability |

The bridge you must not leave unbuilt Security risk and safety risk use different methods and cannot be merged into one matrix. But when a security analysis identifies an exploit path that could lead to patient harm, that harm becomes a hazardous situation in your ISO 14971 risk management file. Two processes; one interface point. Threat models that never reach the risk file are an auditable gap.

Notes — what applies to my product:

PART TWO

Part 2 — Software validation

Verification, validation, assurance

| Term | The question | Compares against |
| --- | --- | --- |
| Verification | Did we build the thing right? | Your own specification |
| Validation | Did we build the right thing? | User needs and clinical reality |
| Assurance | How much evidence is enough? | Risk — what failure would cost |

A test suite with 100% pass and 100% coverage verifies a great deal and validates nothing, if the requirements it traces to were wrong.

Validation is three different jobs

Device software: the software you ship — SaMD, firmware, companion apps, and the cloud services the device depends on to perform its medical function. Design controls and IEC 62304 apply.

Production and quality management system software: MES, test rigs, eQMS, calibration tools, spreadsheets used for release decisions. Validated for intended use; FDA's 2025 Computer Software Assurance guidance applies here.

Infrastructure and tooling: compilers, build pipelines, source control, containers, hosting. Controlled proportionately to what the tool can silently break.

IEC 62304 safety classes

| Class | The software can contribute to | Adds |
| --- | --- | --- |
| A | No injury or damage to health | Planning, requirements, system test, release |
| B | Non-serious injury | Architectural design, integration testing, unit verification |
| C | Death or serious injury | Detailed design to unit level, fullest verification |

Two traps worth naming Classify BEFORE risk controls, on what the software could contribute to absent external risk control measures. Crediting your own mitigations to argue a lower class is circular — and it is exactly what an auditor probes. Segregation must be real. Lower-class processes for lower-class items are legitimate only where the architecture genuinely prevents interference. Separate modules in one binary sharing memory is not segregation, and if the claim fails everything you scoped down is under-documented.

Requirements — the root cause of most failures

Test: could two competent engineers, working independently from this requirement, build things that both pass and behave differently in a clinical setting? If yes, it is not a requirement yet.

| Weak | Stronger |
| --- | --- |
| “The system shall alert the clinician promptly.” | “On detection, the system shall present a visual and audible alert at the clinician workstation within 10 seconds, and shall escalate to the paging interface if unacknowledged within 60 seconds.” |
| “Patient data shall be transmitted securely.” | “All patient data in transit shall be protected using TLS 1.2 or higher with mutual certificate authentication; the device shall reject connections that fail certificate validation and log the rejection.” |

WORKING SESSION 1  ·  12 MINUTES

Classify and scope

The product A remote cardiac monitoring system, intended for ambulatory patients following a cardiac event. · An adhesive wearable patch records single-lead ECG continuously for up to 14 days. · The patch streams over Bluetooth Low Energy to a patient smartphone app, which forwards data to the manufacturer's cloud. · A cloud service runs a machine-learning algorithm that detects and classifies arrhythmias, and flags events for review. · A clinician web portal presents flagged events with the supporting ECG strip for confirmation and diagnosis. · The portal also contains a scheduling and billing module. · Patch firmware can be updated over the air via the app.

1.  Break the product into software functions

For each, decide: SaMD, software in a medical device, or not a device at all?

| Software function | SaMD / SiMD / Not a device | Why |
| --- | --- | --- |

2.  Assign an IEC 62304 safety class to each software item

One sentence of justification you would be willing to defend.

| Software item | Class | Justification |
| --- | --- | --- |

3.  Basic or Enhanced documentation level?

State which criterion drives your answer.

4.  Is this a cyber device under section 524B? Why?

5.  Name three SOUP items you would expect, and what must be documented for each

PART THREE

Part 3 — Cybersecurity

Section 524B — the statutory floor

For a “cyber device”: broadly, one that includes software validated, installed or authorised by the sponsor; has the ability to connect to the internet; and contains technological characteristics that could be vulnerable to cybersecurity threats. Note the ability to connect — not the intention to.

1.  Have a plan to monitor, identify and address postmarket vulnerabilities and exploits, including a coordinated vulnerability disclosure process.

2.  Design, develop and maintain processes providing reasonable assurance the device and related systems are cybersecure, and make available postmarket updates and patches — on a reasonably justified regular cycle, and out-of-cycle for critical vulnerabilities.

3.  Provide a software bill of materials covering commercial, open-source and off-the-shelf components.

4.  Comply with such other requirements as FDA may establish by regulation.

The Secure Product Development Framework

Security risk management: distinct from, but interfaced with, safety risk management. Assess exploitability, not probability.

Security architecture: four views — global system; multi-patient harm; updateability and patchability; security use cases.

Cybersecurity testing: four categories — security requirements testing; threat mitigation testing; vulnerability testing; penetration testing.

Threat modelling — the four questions

| # | Question | Output |
| --- | --- | --- |
| 1 | What are we building? | Data flow diagram with trust boundaries drawn explicitly |
| 2 | What can go wrong? | STRIDE walked across every element and boundary crossing |
| 3 | What are we doing about it? | Controls, traced to requirements and to tests |
| 4 | Did we do a good enough job? | Review, and revisit whenever the architecture changes |

STRIDE: Spoofing · Tampering · Repudiation · Information disclosure · Denial of service · Elevation of privilege. In a clinical setting, denial of service is frequently the highest patient risk — do not import a severity scale that ranks availability third.

The vulnerability management loop

Monitor  →  Triage  →  Assess  →  Decide  →  Act and communicate

CVSS is an input, not an answer: a base score describes a vulnerability in the abstract. It knows nothing about your architecture, your compensating controls, or what happens clinically if the device stops. Use it to prioritise triage; assess clinical consequence separately.

VEX: a machine-readable statement that a component is present but the vulnerability is not exploitable in your product, and why. Without it, SBOM transparency generates noise that exhausts you and your customers.

The question an SBOM exists to answer A serious vulnerability is announced in a widely used component on a Friday afternoon. Your customers' security teams will ask the same thing within hours: are we affected? Without an SBOM that is weeks of archaeology; with a machine-readable one it is a query. Build it as an operational capability — the compliance artefact is a by-product.

WORKING SESSION 2  ·  15 MINUTES  ·  TABLETOP

Monday morning, 07:40

The situation An independent security researcher emails your general support address. A third-party Bluetooth library used in your infusion pump — firmware v3.2, approximately 14,000 units in clinical use — contains an unauthenticated remote code execution flaw. CVSS base score 9.8. They have attached a working proof of concept demonstrated at 10 metres. They have not published, and they ask for a response within 90 days. Two complications: patching requires the pump to be taken out of service for approximately 20 minutes, and roughly 4,200 units are on an earlier hardware revision that cannot accept the update at all.

Round 1 — the first four hours

Who needs to be in the room, and who makes the call?  What three things must you establish before deciding anything?  What do you send back to the researcher today?

Round 2 — day three

Controlled or uncontrolled risk, and on what basis?  What drives your reporting decision, and who do you involve?  What do you tell customers, and when?

Round 3 — day thirty

What are the real options for the 4,200 unpatchable units?  What compensating controls can hospitals apply now?  What would have made this a two-week problem instead?

WORKING SESSION 2  ·  DEBRIEF

What would have made this a two-week problem?

The closing question Everything your table just listed was a decision available to you before Monday morning — an SBOM, a published disclosure route, a rehearsed emergency patch path, universal updateability, a way to reach hospital security teams directly. Capture the list here: it is your Part 4 agenda, and the most useful page in this workbook to take back to your own product.

PART FOUR

Part 4 — Making it work

One lifecycle, not two

| Stage | Validation | Security | Where they meet |
| --- | --- | --- | --- |
| Requirements | SRS from intended use and risk controls | Security requirements from the threat model | One requirements specification |
| Architecture | Architecture, SOUP, segregation | Threat model, trust boundaries, four views | Same diagrams; SOUP seeds the SBOM |
| Risk | ISO 14971 hazard analysis | Exploitability assessment | Separate methods; exploit paths to harm enter the 14971 file |
| Verification | Unit, integration, system testing | Requirements, threat mitigation, vulnerability, penetration testing | One test strategy, one matrix |
| Post-release | Complaints, defects, change control | Vulnerability monitoring, disclosure, patching | One intake, one triage, one change process |

Where to start

| First 30 days — find out where you stand | Days 30–60 — build the capability | Days 60–90 — make it a system |
| --- | --- | --- |
| Determine formally whether each product is a cyber device, and record the reasoning Produce a component inventory for one product, generated from the build Publish a security contact and disclosure policy Name an owner for vulnerability monitoring, with a response time Check SOPs for retired vocabulary — Level of Concern, pre-QMSR citations | Run one threat model on your highest-risk product Audit requirements for testability, especially security requirements Close the traceability gaps you already know about Define the release gate and the emergency patch path — then rehearse it Map cloud and supplier dependencies into change control | Integrate: one requirements specification, one matrix, one intake queue Automate SBOM generation and vulnerability matching in the pipeline Set your patch cadence and be able to justify it Publish end-of-support dates for fielded products Put the security architecture views into your design documentation template |

Five things worth remembering

1.  Security is a safety property, not an IT concern. If an attacker can make your device harm a patient, that is a hazardous situation and it belongs in the risk management file.

2.  The obligations are now lifecycle obligations. The question moved from “was it right at release?” to “can you keep it right?”

3.  You cannot manage what you have not inventoried. SOUP list, SBOM, deployed fleet, component support horizons.

4.  Updateability is the highest-leverage design decision. Every other control defends against threats you have thought of. The ability to ship a fix defends against the ones you have not.

5.  Integrate, or pay twice. One requirements specification, one traceability matrix, one intake queue, one change process.

SELF-ASSESSMENT

Readiness self-assessment

Score each statement: 2 = true and evidenced,  1 = partly true,  0 = not true. Answer for your lead product, not for the company as a whole.

| # | Statement | Score |
| --- | --- | --- |
| 1 | We have a written determination of whether this product is a cyber device under section 524B. |  |
| 2 | We can produce a complete, versioned inventory of third-party software components, generated from the build. |  |
| 3 | Every security requirement lives in the software requirements specification and is traced to a test. |  |
| 4 | We have a current threat model, produced at architecture time and revisited when the architecture changed. |  |
| 5 | Our traceability matrix is maintained continuously, not reconstructed before submissions. |  |
| 6 | Risk control effectiveness — not just implementation — is verified and evidenced. |  |
| 7 | There is a published route for an outside researcher to report a vulnerability, and a named owner who monitors it. |  |
| 8 | We monitor vulnerability feeds against our component inventory, and every match gets a recorded disposition. |  |
| 9 | We have an emergency patch path that is defined, pre-validated, and has been exercised at least once. |  |
| 10 | Every fielded product can receive an authenticated, integrity-checked update, with rollback. |  |
| 11 | We have published end-of-support dates for products currently in clinical use. |  |
| 12 | Security findings with a patient-harm pathway are carried into the ISO 14971 risk management file. |  |

| Total | Where that puts you | What to do first |
| --- | --- | --- |
| 0–8 | Level 1 — Reactive. Documentation is produced for submissions; no inventory of what is deployed. | Start with the 30-day list. Inventory and a security contact, in that order. |
| 9–15 | Level 2 — Defined. Procedures exist and are followed, but records are largely produced by hand. | Threat model one product. Move security requirements into the SRS. |
| 16–20 | Level 3 — Integrated. One lifecycle; traceability maintained live. | Automate. SBOM generation and vulnerability matching in the pipeline. |
| 21–24 | Level 4–5 — Automated to anticipatory. Evidence generated by the pipeline; designing for year ten. | Crypto-agility, component support horizons, planned end of support. |

Reading your score honestly Score this twice — once for validation maturity and once for security maturity. The gap between the two is usually the most useful thing you learn. Established manufacturers are often strong on validation and weak on security; newer digital-health companies are frequently the reverse. The remedy differs: if your validation system is mature, you do not need a new system, you need to extend the one you have.

REFERENCE

Documentation level and safety class

FDA documentation level

Enhanced — if ANY of these is true · The device is a constituent part of a combination product. · The device is intended to test blood donations for transfusion-transmitted infections, or to determine donor/recipient compatibility; or it is a blood establishment computer software product. · The device is Class III. · A failure or latent flaw in the device software function could present a probable risk of death or serious injury — to a patient, a user, or anyone else — assessed BEFORE risk control measures are applied.

Basic — everything else Basic is the default, not a concession to be argued for. It still requires a genuine documentation package. Enhanced adds depth rather than new categories: fuller design documentation, more detailed V&V evidence including unit and integration results, and a fuller configuration management and maintenance picture. The level applies to the device software function. One submission can contain functions at different levels.

Basic/Enhanced and IEC 62304 Class A/B/C are independent frameworks answering different questions. They correlate; they do not map. Record both decisions with independent reasoning — a reviewer who sees “Enhanced because Class C” has learned that you did the analysis once.

The premarket software documentation package

| Element | What makes it adequate — not merely present |
| --- | --- |
| Software description and device architecture | A reviewer with no prior knowledge can understand what it does, where it runs and what it talks to. Include cloud and mobile components. |
| Risk management file | Hazards, hazardous situations and harms, with risk controls traced to requirements and to the verification that proves each control works. |
| Software requirements specification | Testable, unambiguous, complete — including performance, interface, security and alarm requirements. |
| Architecture and design documentation | Modules, interfaces, data flows, SOUP items and their integration. |
| Development environment description | Lifecycle plan, configuration management and the maintenance process — including how defects and patches are handled after release. |
| V&V documentation | Unit, integration and system level testing traceable to requirements. Enhanced expects the underlying evidence, not only a summary. |
| Revision history and unresolved anomalies | Every open defect with a documented clinical impact assessment. Declaring nothing open is a red flag, not a clean bill of health. |

REFERENCE

Common findings that recur

Recurring themes in software-related observations. None of these are exotic — they are all process discipline.

Validation not completed before release, or performed retrospectively.

Requirements too vague to be testable.

Traceability gaps — requirements with no test, tests with no requirement, risk controls never verified for effectiveness.

Changes released with no documented impact assessment.

Known anomalies carried forward with no clinical impact evaluation.

SOUP unmanaged, and security requirements that live outside the SRS and so are never tested.

The pattern: these are failures to maintain something, not to create it. The initial build is usually fine; the lifecycle is where programmes erode.

The one to act on first Pick the finding above you would least like an investigator to ask about tomorrow. That is your first agenda item on Monday — carry it to the action plan at the back of this workbook.

Which one, and what will you do about it?

REFERENCE

Security baseline checklist

Control categories to design against

Design each against your threat model, not against this list. A control you cannot justify from a specific threat is cost without security — and a threat with no control is the finding.

| Category | What good looks like | ✓ |
| --- | --- | --- |
| Authentication | Identity established for users, services and connected components. No shared or default credentials; no hardcoded secrets. |  |
| Authorisation | Least privilege enforced server-side, not in the UI. Clinical, administrative and service roles separated. |  |
| Cryptography | Current algorithms, correct implementations, real key management. Keys rotatable across the service life. |  |
| Code and data integrity | Signed firmware, verified boot, protected configuration and calibration data. The device refuses code it cannot authenticate. |  |
| Confidentiality | Data protected in transit, at rest, and at end of life — including diagnostic logs and memory. |  |
| Event detection and logging | Security-relevant events recorded, retained and retrievable. An incident can be reconstructed. |  |
| Resiliency and recovery | Clinical function maintained or safely degraded under attack; recovery to a known-good state. |  |
| Secure updates | Authenticated, integrity-checked, rollback-capable — including behaviour on an interrupted update. |  |

The four security architecture views

| View | The question it answers |
| --- | --- |
| Global system | What is the whole picture — the device, everything it connects to, every interface and data flow, including cloud services, mobile apps, service tools and update infrastructure? If it is not on this diagram, it was not assessed. |
| Multi-patient harm | How does a compromise propagate beyond one device? What is the largest number of patients one successful attack could affect? Anything shared fleet-wide — a signing key, a default credential, a common server — is a pathway. |
| Updateability and patchability | How does software get into the device, and how would you deploy an emergency fix? Authentication, integrity verification, rollback, and behaviour on interruption. |
| Security use cases | Walk scenarios end to end — clinician login, firmware update, service engineer connection, remote support, decommissioning. Controls fail in sequence, not in isolation. |

SBOM — minimum content

| Field | Note |
| --- | --- |
| Supplier name | Who produced the component |
| Component name and version | “Latest” is not a version |
| Unique identifiers | For automated matching against vulnerability feeds |
| Dependency relationships | Including transitive dependencies — this is where URGENT/11 lived |
| Author of the SBOM data and timestamp | Provenance of the inventory itself |
| Level of support from the supplier | Expected by FDA in addition to the baseline fields |
| End-of-support and end-of-life dates | Expected by FDA. Compare against your device's service life. |

Machine-readable — SPDX or CycloneDX. A PDF of a table is not an SBOM: the entire value is automated matching. Generate it from the build, not from a spreadsheet someone maintains.

Postmarket — the four obligations that never close

Coordinated vulnerability disclosure: a published route for researchers to reach you, a monitored address, and a stated policy on scope and timelines. ISO/IEC 29147 and 30111 are the references. Researchers who cannot reach you publish anyway.

Controlled vs uncontrolled risk: controlled means residual risk to health is acceptable; uncontrolled means it is not, and must be remediated. Unlike safety risk, uncontrolled security risk is not something you accept against benefit.

Reporting thresholds: routine updates that reduce risk to an acceptable level are treated differently from corrections reportable under 21 CFR 806. Read FDA's 2016 postmarket guidance before you need it, and involve regulatory on day one.

Ecosystem participation: information sharing through an ISAO, engagement with CISA advisories, and relationships with the researcher community.

REFERENCE

Glossary

| Term | Meaning |
| --- | --- |
| 524B | Section of the US Federal Food, Drug, and Cosmetic Act, added in 2023, imposing statutory cybersecurity requirements on sponsors of “cyber devices.” |
| AAMI SW96 | Standard for security risk management by medical device manufacturers; the security counterpart to ISO 14971's safety process. |
| Basic / Enhanced documentation level | FDA's two-tier framework for how much software documentation a premarket submission requires. Replaced Level of Concern in 2023. |
| CDS | Clinical decision support. Certain CDS software is excluded from the device definition by the 21st Century Cures Act, on narrow conditions including that the clinician can independently review the basis for the recommendation. |
| Controlled / uncontrolled risk | FDA postmarket cybersecurity vocabulary. Controlled: residual risk to health is acceptable. Uncontrolled: it is not, and remediation is expected. |
| CVD | Coordinated vulnerability disclosure — the process by which outside researchers report vulnerabilities and manufacturers respond. ISO/IEC 29147 and 30111 refer. |
| CVSS | Common Vulnerability Scoring System. A context-free severity score; an input to triage, not a statement of clinical risk. |
| Cyber device | Under 524B, broadly: a device that includes software validated, installed or authorised by the sponsor; has the ability to connect to the internet; and has technological characteristics that could be vulnerable to cybersecurity threats. |
| Design controls | The systematic design process required by the quality system regulation — planning, inputs, outputs, review, verification, validation, transfer, changes, and the design history file. |
| DHF | Design history file. The compilation of records describing the design history of a finished device. |
| Exploitability | The feasibility of an attack given required access, skill, equipment and window of opportunity. Used in security risk assessment in place of probability. |
| GPSV | FDA's General Principles of Software Validation (2002). Still the foundational FDA statement of what software validation means. |
| IEC 62304 | The medical device software lifecycle standard. Defines processes and outputs; method-agnostic. Safety classes A, B and C. |
| IEC 81001-5-1 | Security activities in the health software product lifecycle. Operationalises security within a 62304-shaped lifecycle. |
| IMDRF | International Medical Device Regulators Forum. Publisher of the SaMD framework documents (N10, N12, N23, N41) and guidance on cybersecurity and legacy devices. |
| ISAO | Information Sharing and Analysis Organization. Participation is explicitly contemplated in FDA's postmarket cybersecurity guidance. |
| ISO 13485 | Quality management systems for medical devices. Incorporated by reference into 21 CFR 820 under the QMSR. |
| ISO 14971 | Application of risk management to medical devices. The safety risk process; ISO/TR 24971 provides guidance on applying it. |
| Level of Concern | Retired FDA framework (Minor / Moderate / Major) superseded in 2023 by Basic and Enhanced documentation levels. If your templates still use it, they are out of date. |
| MDS2 | Manufacturer Disclosure Statement for Medical Device Security. The established vehicle for disclosing security characteristics to purchasers. |
| PCCP | Predetermined Change Control Plan. Pre-authorisation of specified future changes, their validation protocol and their impact assessment. |
| QMSR | FDA's Quality Management System Regulation. Amends 21 CFR 820 to incorporate ISO 13485:2016; in force from February 2026. |
| SaMD | Software as a Medical Device — software intended for a medical purpose that performs that purpose without being part of a hardware medical device. |
| SBOM | Software bill of materials. A machine-readable inventory of software components and their relationships. Required under 524B. |
| SiMD | Software in a medical device — software embedded in, or driving, a hardware medical device. |
| SOUP | Software of unknown provenance. Software you did not develop and cannot fully characterise, but remain responsible for. |
| SPDF | Secure Product Development Framework. FDA's organising concept: security risk management, security architecture, and cybersecurity testing. |
| SPDX / CycloneDX | The two widely used machine-readable SBOM formats. |
| STRIDE | A threat classification mnemonic: Spoofing, Tampering, Repudiation, Information disclosure, Denial of service, Elevation of privilege. |
| Trust boundary | A point in a system where data or control crosses between components with different levels of trust. Where threat modelling concentrates. |
| VEX | Vulnerability Exploitability eXchange. A machine-readable statement that a component is present but the vulnerability is not exploitable in your product, and why. |

REFERENCE

Standards and guidance — and what each is for

Dates are stated as at preparation. Confirm current versions before relying on any of these, particularly for a submission.

FDA guidance

| Document | Read it for |
| --- | --- |
| Content of Premarket Submissions for Device Software Functions (June 2023) | What goes in the submission, and the Basic / Enhanced documentation level framework. Start here. |
| Cybersecurity in Medical Devices: Quality Management System Considerations and Content of Premarket Submissions (June 2025) | The SPDF, security architecture views, SBOM expectations, and the four categories of security testing. Supersedes the September 2023 edition and adds a section on s.524B. |
| Postmarket Management of Cybersecurity in Medical Devices (Dec 2016) | Controlled vs uncontrolled risk, and the circumstances bearing on reporting obligations. Read before you need it. |
| General Principles of Software Validation (Jan 2002) | The foundational statement of what validation means. Old, not withdrawn. |
| Off-The-Shelf Software Use in Medical Devices; Deciding When to Submit a 510(k) for a Software Change (Oct 2017) | Obligations for third-party software you incorporate, and the decision logic for whether a change needs a new submission. |
| Marketing Submission Recommendations for a PCCP for AI-Enabled Device Software Functions (final, Dec 2024) | Pre-authorising future changes: description, modification protocol, impact assessment. |
| Computer Software Assurance for Production and Quality Management System Software (final, Sept 2025) | Risk-based assurance for production and QMS software. Supersedes Section 6 of the General Principles of Software Validation. |
| Clinical Decision Support Software (Sept 2022) | Where the Cures Act CDS exclusion begins and ends. |

Standards and technical reports

| Document | Read it for |
| --- | --- |
| IEC 62304 | The software lifecycle: processes, outputs, safety classes, SOUP obligations, maintenance. |
| IEC 81001-5-1 | How to put security activities inside a 62304-shaped lifecycle rather than alongside it. The best implementation reference. |
| ISO 14971 + ISO/TR 24971 | Safety risk management, and practical guidance on applying it. |
| ISO 13485 | The QMS, now incorporated into 21 CFR 820 via the QMSR. |
| AAMI SW96 | Security risk management as a distinct process that interfaces with 14971. |
| AAMI TIR57 / TIR97 / TIR45 | Security risk management principles, postmarket security management, and agile practice within a regulated lifecycle. |
| ISO/IEC 29147 / 30111 | Vulnerability disclosure and vulnerability handling processes. |

International and community resources

| Document | Read it for |
| --- | --- |
| IMDRF SaMD series — N10, N12, N23, N41 | Definitions, risk categorisation, QMS and clinical evaluation for SaMD. |
| IMDRF N60 (cybersecurity) and N73 (legacy) | International principles, and what is expected for devices already in the field. |
| MITRE / MDIC Medical Device Threat Modeling Playbook | Free and genuinely practical. The best first document for a team that has never run a threat model. |
| MITRE / MDIC rubric for applying CVSS to medical devices | Why generic severity scoring misleads in a clinical context, and what to do instead. |
| MDCG 2019-11 and 2019-16; CISA advisories | EU software qualification and classification, EU cybersecurity guidance, and live advisories affecting medical and industrial devices. |

THE PAGE THAT MATTERS

Your action plan

Fill in at least three rows before you leave the room. A commitment made here is the only part of a seminar that reliably survives the journey back to the office.

| What I will do | Who owns it | By when |
| --- | --- | --- |

Ten questions to ask your team on Monday

1.  Is this product a cyber device under 524B — and where is that determination written down?

2.  Could we produce a complete, versioned component inventory for it this week?

3.  Who receives a vulnerability report from outside, and what happens in the first hour?

4.  What is our device's IEC 62304 safety class, and would the reasoning survive a challenge?

5.  Do our templates still say “Level of Concern” or cite 21 CFR 820 clauses superseded by the QMSR?

6.  Where do our security requirements live — and are they traced to tests?

7.  If a cloud provider deprecated a runtime with 90 days' notice, whose change control handles it?

8.  Can every fielded unit accept an authenticated update, and have we ever exercised the emergency path?

9.  What is our end-of-support date, and have we told our customers?

10.  Where are validation and security duplicating each other's work right now?

If you do only one thing Publish a security contact and a disclosure policy, and name the person who monitors it. It is days of work, it costs almost nothing, it is a statutory expectation under 524B, and it changes what happens on the day a researcher finds something in your product.