QUICK REFERENCE  ·  SIDE 1  ·  SOFTWARE VALIDATION

Scoping and classifying

Four independent decisions. Make each on its own terms and write down the reasoning for each.

1.  Is it a device at all?

| SaMD | SiMD | Not a device |
| --- | --- | --- |
| Software intended for a medical purpose that performs that purpose WITHOUT being part of a hardware device. Algorithm, diagnostic platform, dosing app. | Software embedded in, or driving, a hardware medical device. Pump firmware, ventilator control, imaging acquisition. | Excluded by the 21st Century Cures Act: administrative support, general wellness, record-keeping, and certain CDS where the clinician can independently review the basis. |

Classify at the level of the software FUNCTION, not the product. One product routinely contains all three.

2.  IEC 62304 safety class

| Class | Can contribute to | Adds to the process |
| --- | --- | --- |
| A | No injury or damage to health | Planning, requirements, system test, release, plus the continuous processes |
| B | Non-serious injury | Architectural design, integration testing, unit verification |
| C | Death or serious injury | Detailed design to unit level, fullest verification and documentation |

Classify BEFORE risk controls. Crediting your own mitigations to argue a lower class is circular. Segregation to a lower class is legitimate only where the architecture genuinely prevents interference — and is documented and verified.

3.  FDA documentation level

| ENHANCED — if ANY is true | BASIC — everything else |
| --- | --- |
| · Constituent part of a combination product · Intended to test blood donations for transfusion-transmitted infections or determine donor/recipient compatibility; or blood establishment computer software · Class III device · A failure or latent flaw could present a probable risk of death or serious injury — to a patient, user or anyone else — assessed BEFORE risk controls | Basic is the default, not a concession. It still requires a real package: description, risk file, requirements, architecture, development environment, V&V summary, unresolved anomalies. Enhanced adds depth, not new categories. The level applies per software function — one submission can contain both. |

4.  Is it a cyber device under 524B?

Broadly: includes software validated, installed or authorised by the sponsor  +  has the ABILITY to connect to the internet  +  has technological characteristics that could be vulnerable. Ability, not intention — a service port or a disabled radio still counts.

Requirements test

Could two competent engineers, working independently from this requirement, build things that both pass and behave differently in a clinical setting? If yes, it is not a requirement yet. “Promptly”, “securely”, “reliably” and “robustly” are not requirements — decompose or delete.

Traceability chain

User need  →  System requirement  →  Software requirement  →  Design element  →  Test case  →  Result

Forwards: every user need reaches a verified test.  Backwards: every test and every piece of functionality traces to a requirement — orphans are undocumented features, which are undocumented risk. Risk controls attach to the chain: verify implementation AND effectiveness, separately.

QUICK REFERENCE  ·  SIDE 2  ·  CYBERSECURITY

Securing and maintaining

Section 524B — four statutory obligations

| 1 | Plan to monitor, identify and address postmarket vulnerabilities and exploits, including coordinated vulnerability disclosure |
| --- | --- |
| 2 | Processes giving reasonable assurance the device and related systems are cybersecure; make updates and patches available on a reasonably justified regular cycle, and out-of-cycle for critical vulnerabilities |
| 3 | Provide an SBOM covering commercial, open-source and off-the-shelf components |
| 4 | Comply with such other requirements as FDA may establish by regulation |

Threat modelling starter — four questions, one afternoon

| # | Question | Output |
| --- | --- | --- |
| 1 | What are we building? | Data flow diagram with trust boundaries drawn explicitly |
| 2 | What can go wrong? | STRIDE walked across every element and every boundary crossing |
| 3 | What are we doing about it? | Controls, traced to requirements and to tests |
| 4 | Did we do a good enough job? | Review, and revisit whenever the architecture changes |

STRIDE:  Spoofing · Tampering · Repudiation · Information disclosure · Denial of service · Elevation of privilege.  In clinical settings denial of service is often the highest patient risk.

Security control baseline

| Authentication | No shared or default credentials; no hardcoded secrets | Confidentiality | In transit, at rest, and at end of life — including logs |
| --- | --- | --- | --- |
| Authorisation | Least privilege, enforced server-side not in the UI | Detection & logging | Security events recorded, retained, retrievable |
| Cryptography | Current algorithms; keys rotatable across service life | Resiliency & recovery | Maintain or safely degrade; recover to known-good |
| Code & data integrity | Signed firmware, verified boot, protected config | Secure updates | Authenticated, integrity-checked, rollback-capable |

Design each against your threat model. A threat with no control is the finding; a control with no threat is cost without security.

The four architecture views

| Global system | Everything it connects to — including your own update servers, telemetry and service tools |
| --- | --- |
| Multi-patient harm | Largest number of patients one attack could affect? Anything shared fleet-wide is a pathway. |
| Updateability | How would you ship an emergency fix? What if an update is interrupted at 60%? |
| Security use cases | Login, update, service access, remote support, decommissioning — controls fail in sequence |

SBOM minimum content

Supplier  ·  component name  ·  version  ·  unique identifiers  ·  dependency relationships  ·  author of the SBOM data  ·  timestamp  —  plus, expected by FDA:  level of supplier support  ·  end-of-support and end-of-life dates.

Machine-readable (SPDX or CycloneDX), generated from the build, including transitive dependencies. A PDF of a table is not an SBOM.

Vulnerability response — the loop

Monitor  →  Triage (present, reachable, exploitable in YOUR product?)  →  Assess (CVSS as an input; clinical consequence as the output)  →  Decide (controlled or uncontrolled? patch now, on cycle, compensating control, or no action)  →  Act and communicate (deploy, tell customers, publish VEX).  Every match gets a recorded disposition — including “not exploitable”, with the reason.

Four kinds of security testing

Requirements — does each security requirement behave as specified, including correctly refusing?   Threat mitigation — do the controls defeat the specific attacks in your threat model, in the assembled system?   Vulnerability — abuse cases, malformed input, fuzzing, attack surface analysis, chaining, SAST/DAST, known-vuln scans.  Penetration — skilled adversarial testing, defined scope, every finding dispositioned.

A penetration test report with no remediation record is worse than no test: you documented that you knew.