MEDICAL DEVICE SEMINAR  ·  WORKSHOP

Hands-On Exercise Pack

Risk Management Workshop (ISO 14971): three exercises, 35 minutes

Participant worksheet  ·  Elder Consulting, LLC  ·  September 2026

Write on this. Exercise A on pages 2 to 3, Exercise B on pages 4 to 5, Exercise C on page 6.

Device for all three exercises: a programmable ambulatory infusion pump for home use, with a single-use administration set. The same device as the Excel toolkit.

### The device, in one paragraph

A battery-powered programmable infusion pump, prescribed for home use and operated by the patient or a lay carer. It delivers a titrated therapy through a single-use administration set over periods of hours to days. It has a drug library with configurable dose limits, updated over a wireless network; an air-in-line detector; an occlusion alarm; and a rechargeable battery with a remaining-time indicator. The therapy is life-sustaining for the indicated population, and the alternative is inpatient infusion.

You do not need to know this device. Everything you need is on the page. If a detail matters to your answer and it is not stated, write down the assumption you are making - that is part of the exercise.

# Exercise A  ·  Repair the hazard chain

15 minutes: 10 in groups of three or four, 5 taking answers.

### The row, exactly as it appears in the file

| Column | What the row says |
| --- | --- |
| ID | HZ-15 |
| Device function or part | Pumping mechanism |
| Hazard | Pump failure |
| Foreseeable sequence of events | Pump fails. |
| Hazardous situation | Pump failure. |
| Harm | Harm to patient. |
| MDR Annex I GSPR link | (blank) |
| P1 / P2 / Severity | 2 / 2 / 3 |
| Basis for the estimate | (blank) |
| Initial risk | Acceptable |
| Risk control | Mitigated by design |
| Verification of implementation | (blank) |
| Verification of effectiveness | (blank) |
| Residual P1 / P2 / Severity | (blank) |
| Status | Open |

### Step 1  ·  Take the row apart

For each of the four positions in the chain, write what the row has actually put there, and say whether it is empty, duplicated, or in the wrong position.

| Position | What the row actually says | Problem |
| --- | --- | --- |
| Hazard (3.4) |  |  |
| Foreseeable sequence of events (5.4) |  |  |
| Hazardous situation (3.5) |  |  |
| Harm (3.3) |  |  |

### Step 2  ·  Rebuild it

Write as many rows as this hazard needs - one per hazardous situation. For each, assign P1, P2 and severity, and write the basis you would cite for each probability. Use extra paper if you need more than three rows.

### Exercise A, step 2  ·  Rebuild grid

Rows 1 and 2 below; row 3 and step 3 are on the next page.

Row 1

| Field | Your answer |
| --- | --- |
| Hazard |  |
| Foreseeable sequence of events |  |
| Hazardous situation |  |
| Harm |  |
| P1 and its basis |  |
| P2 and its basis |  |
| Severity |  |
| MDR Annex I link |  |

Row 2

| Field | Your answer |
| --- | --- |
| Hazard |  |
| Foreseeable sequence of events |  |
| Hazardous situation |  |
| Harm |  |
| P1 and its basis |  |
| P2 and its basis |  |
| Severity |  |
| MDR Annex I link |  |

Row 3

| Field | Your answer |
| --- | --- |
| Hazard |  |
| Foreseeable sequence of events |  |
| Hazardous situation |  |
| Harm |  |
| P1 and its basis |  |
| P2 and its basis |  |
| Severity |  |
| MDR Annex I link |  |

### Step 3  ·  Write the finding

Draft the deficiency a notified body reviewer would raise against the original row. Cite the clause. One or two sentences.

# Exercise B  ·  Red-team the file

14 minutes: 9 in groups, 5 taking answers, one finding per table with no repeats.

You are the notified body technical reviewer assessing this technical documentation. You have the excerpt below and nothing else. There are at least six defensible findings. Write them, cite the clause or the MDR Annex I point, then rank them: which one would you raise as a major nonconformity, and why?

### The excerpt

| Field | HZ-16 | HZ-11 |
| --- | --- | --- |
| Hazardous situation | Alarm silenced. | Residue and microbial load transferred between patients |
| Harm | Delay of therapy. | Cross-infection |
| MDR Annex I link | 1 | 1, 11.1, 11.2 |
| Initial P1 / P2 / S | 5 / 5 / 4 | 3 / 2 / 3 |
| Basis for the estimate | Engineering judgement. | Cleaning validation CV-03; material compatibility MC-11 over 200 cycles |
| Initial risk | Not acceptable | Acceptable |
| Risk control | Warning added to IFU section 6 telling the user not to silence the alarm without checking the line | Validated cleaning agent list and procedure in the IFU; compatibility-tested enclosure material; label symbol pointing to the cleaning section |
| Control category | 3 (information for safety) | 3 (information for safety) |
| Why a higher-priority option was not practicable | (blank) | A fully sealed enclosure was rejected on thermal grounds; a single-patient-use durable was rejected on cost of therapy |
| Residual P1 / P2 / S | 1 / 1 / 4 | 3 / 2 / 3 |
| Residual risk | Acceptable | Acceptable |
| Verification of implementation | IFU approval record DOC-2230 | IFU approval DOC-2214; label artwork approval LA-118 |
| Verification of effectiveness | (blank) | CV-03 cleaning validation with the listed agents over 200 cycles; MC-11 material compatibility |
| Residual risk disclosure | (blank) | IFU section 7.3; label symbol |
| New risks assessed (7.5) | (blank) | RA-NEW-11 (an unlisted agent used anyway) |
| Status | Closed | Closed |

### Exercise B  ·  Findings grid

| # | Finding | Clause or GSPR | Major? |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |
| 8 |  |  |  |

### Which one is the major nonconformity, and why?

### Is there anything in this excerpt that is done well?

# Exercise C  ·  Your 30-day plan

8 minutes: 5 writing on your own, 3 sharing first actions and Q&A.

Not "what should we improve" but "what will I have done by this day next month". One action per row. An owner who is in the building. A date inside 30 days.

| # | What I will change | Clause or GSPR it answers | Owner | Done by | How I will know it worked |
| --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |
| 5 |  |  |  |  |  |

### If you need prompts: six checks from today

- Does the risk management plan state what you do when the probability of occurrence of harm cannot be estimated? (Clause 4.4 d)

- Does every risk control have a verification of EFFECTIVENESS record, distinct from its implementation record? (Clause 7.2)

- Does any justification for skipping a higher-priority control option rest on cost? (MDR Annex I point 2)

- Does the procedure have a Clause 7.6 completeness step, applied to the acceptable rows too?

- Can you find the risk file revision that one old complaint produced? (Clauses 10.3 and 10.4)

- Do the acceptability criteria in the risk management plan and the benefit-risk acceptability parameters in the clinical evaluation plan agree?

### Notes

Prepared by Elder Consulting, LLC. Workshop material; not legal or regulatory advice.