MEDICAL DEVICE SEMINAR  ·  CASE STUDY WORKSHOP PACK

ISO 13485 & Audit Readiness

Four composite cases on surviving unannounced audits and maintaining a global QMS

Presented by: Elder Consulting, LLC   ·   Table number: ______   ·   Date: ____________

Companion document: Participant Handout

# How to use this pack

The four cases are composites built from real audits and inspections. Names, places, products and details have been changed; the mechanics have not. Each case runs for about seven minutes: two minutes to read, three to discuss at your table, and two for a plenary debrief. The exhibits are short excerpts from the kind of documents the auditors saw. Read them as an auditor would.

- There is no single right answer; there are answers that hold up in an audit and answers that do not.

- Appoint a spokesperson for each case so the debrief moves quickly.

- At the end of every case, write down one thing you would change at your own site. That list is the real output of the session.

### Grading vocabulary you will need

| Regime | How findings are expressed |
| --- | --- |
| Notified body (EU MDR) | Minor and major nonconformities; certificate suspension or withdrawal for refusal of access or systemic failures |
| FDA | Form FDA 483 observations; inspection classified NAI / VAI / OAI; Warning Letter if the response is inadequate |
| MDSAP | Grades 1–5: direct-impact clauses start at 3, indirect at 1; +1 for no documented process; +1 for a repeat; grades 4–5 notified to all regulators within 5 business days |

# Case 1: The Tuesday morning knock

## Background

- Legal manufacturer in Germany; Class IIb powered surgical instruments.

- About 220 staff on a single manufacturing site; MDR certificate issued by a notified body.

- The QA Director is on leave for two weeks. The reception instructions name no deputy.

## The day

| Time | What happened |
| --- | --- |
| 08:10 | Two notified body auditors present credentials at reception and state they are conducting an unannounced audit under MDR Annex IX. Reception calls HR, then the QA Director's mobile. Voicemail. |
| 09:00 | The Production Manager, alerted by a colleague, takes the auditors to a meeting room. The 50-minute delay is written into the audit report. |
| 09:15 | Opening meeting. The auditors state they will sample from final assembly and trace the sampled units through the QMS. |
| 09:40 | Three finished units are selected. The auditors request the DHRs, the current DMR and the calibration status of the torque tools on the line. |
| 10:30 | One torque driver is three weeks past its calibration due date and still in use. No impact assessment exists. |
| 11:15 | The auditors ask for the validation of an automated dispensing station installed four months earlier. IQ/OQ/PQ exists; the change record has no design or risk impact assessment. |
| 13:40 | PCB supplier file: approved; last audit four years ago; the quality agreement is silent on notified body access and change notification. |

## Exhibits

EXHIBIT 1A Reception desk instruction (excerpt) Visitors without an appointment: ask the visitor to wait in the lobby and contact the host department. Auditors and inspectors: contact the QA Director (ext. 210).

EXHIBIT 1B Calibration register (excerpt, audit date 8 September) Asset TQ-017 ¦ Torque driver 0.5–5 Nm ¦ Final assembly, station 3 Last calibration: 14 May ¦ Due: 14 August ¦ Status: IN USE Recall action: none recorded

EXHIBIT 1C Change record CR-24-118 (excerpt) Description: install automated adhesive dispensing station replacing manual syringe dispensing at station 5. Validation: IQ/OQ/PQ per VAL-118 - complete. Design impact assessment: [blank] Risk management file update: [blank] Notified body notification assessed: [blank]

## Discuss at your table

- What should the first 30 minutes have looked like? Write the sequence of calls and actions.

- Rank the three findings from most to least serious, and explain your ranking. Which would you expect the notified body to grade as major?

- What did the 50-minute delay tell the auditors before they saw a single record?

- Exhibit 1C: who should have been required to sign the blank fields, and what would have happened to the change if they had?

### Table notes

### One thing I will change at my site

# Case 2: The supplier they visited first

## Background

- Legal manufacturer in the Netherlands; Class IIa sterile single-use devices; MDR Annex IX certificate.

- Final assembly, packaging and sterilization release are outsourced to a contract manufacturer in Penang, Malaysia, named in the technical documentation as a critical subcontractor.

- The device is manufactured in campaigns of two to three weeks, several times a year.

## The day

| Time | What happened |
| --- | --- |
| 09:00 (Penang) | Two notified body auditors arrive unannounced at the subcontractor's gate. Security refuses entry: no appointment, no visitor request in the system. |
| 09:20 | The subcontractor's QA manager calls the manufacturer's RA lead in the Netherlands. It is 03:20 there. No answer. |
| 12:05 | Contact is made with the manufacturer at 08:00 CET. Access is authorized. Three hours of delay are recorded in the report. |
| 12:30 | The manufacturer's device is not in production that week. The auditors sample from finished-goods stock and review the DHRs of the last two campaigns. |
| 14:15 | Four deviations closed locally in the last six months; none was reported to the manufacturer. The quality agreement requires notification within five days. |
| 15:00 | One deviation is a seal-temperature excursion on a pouch sealer, dispositioned "use as is" by the subcontractor without the manufacturer's involvement. |
| 16:30 | Closing meeting by video with the manufacturer. The manufacturer's supplier scorecard, shown on request, rates the subcontractor "green" for the entire period. |

## Exhibits

EXHIBIT 2A Quality agreement, clause 7.2 (as signed) The Supplier shall permit the Manufacturer to audit the Supplier's facilities upon thirty (30) days' written notice, not more than once per calendar year.

EXHIBIT 2B Supplier scorecard, Q1-Q2 (manufacturer's record) On-time delivery: 98% - GREEN Incoming inspection rejects: 0.2% - GREEN Deviations reported to manufacturer: 0 - GREEN Complaints attributable to supplier: 0 - GREEN

EXHIBIT 2C Subcontractor deviation DEV-0231 (excerpt) Event: seal temperature recorded at 168 C against a validated range of 175-185 C for approximately 40 minutes during lot 24K07. Investigation: seal strength samples (n=5) from the affected period passed. Disposition: use as is. Approved: Production Supervisor; QA Officer (site). Customer notification: not required.

## Discuss at your table

- Whose audit was this, really, and who received the findings?

- What does ISO 13485 §4.1.5 (control of outsourced processes) require of the manufacturer here, and where exactly did the manufacturer's control fail?

- Exhibit 2B: why did a scorecard with four green cells hide the problem? What would you measure instead?

- How should production calendars and site access be arranged with the notified body for a campaign-manufactured device?

- Rewrite clause 7.2 of the quality agreement in two or three sentences.

### Table notes

### One thing I will change at my site

# Case 3: The first inspection under the QMSR

## Background

- US manufacturer in Indiana; Class II reusable orthopedic instruments and sterilization trays; about 400 staff.

- ISO 13485 certified for ten years; MDSAP certificate for the Canadian market. Last FDA inspection 2021, classified NAI.

- The QMSR gap assessment completed in 2025 concluded: "We are ISO 13485 certified; no procedural changes required."

## The day

| Time | What happened |
| --- | --- |
| 09:00 | An FDA investigator presents credentials and Form FDA 482. The inspection follows the QMSR structure, not QSIT. |
| 09:20 | The investigator asks for the procedure index cross-referenced to ISO 13485 clauses and to the QMSR. The firm has an ISO index only. |
| 09:45 | Request for the last two management reviews and the internal audit schedule and results. The firm's readiness procedure still cites the retired 820.180(c) exemption. Forty minutes pass in calls with corporate RA before the records are provided. |
| 11:00 | Management review minutes contain no regulatory-reporting metrics; two actions from 2024 are still open. |
| Day 2 | Complaints: 6 of 20 sampled files have no documented MDR reportability decision; complaint records do not capture the UDI (820.35(a)). |
| Day 2 | Labeling: label and IFU inspection records for one tray family show no verification of label content (820.45). |
| Day 3 | CAPA: two CAPAs for the same cleaning-validation issue, both "closed - effective" with no effectiveness data. |

## Exhibits

EXHIBIT 3A Inspection readiness procedure QP-015 rev C (excerpt, effective 2019) Per 21 CFR 820.180(c), records of management review, internal quality audits and supplier audits are exempt from FDA review. The FDA investigator shall be informed that such audits have been performed and shall be shown the certification of completion only.

EXHIBIT 3B Complaint record CMP-25-0417 (fields as completed) Product: tray system model T-40 ¦ Lot / UDI: [blank] Complaint: "Instrument tray latch failed to close during reprocessing; tray sent for service." Investigation: latch spring fatigue; spring replaced under service. MDR reportability decision: [blank] ¦ Rationale: [blank] Reply to complainant: yes, 14 March.

EXHIBIT 3C CAPA-24-031 closure statement Root cause: operator did not follow cleaning validation protocol. Corrective action: retraining completed 12 November 2024. Effectiveness: verified - no recurrence observed at closure. Closure date: 15 November 2024.

## Discuss at your table

- What mindset caused the 40-minute standoff over the management review records, and what did it cost the firm?

- Which of the eventual observations are QMSR-specific, and which are plain ISO 13485 failures the firm was already certified against?

- Exhibit 3C: list everything wrong with this closure statement. What would an acceptable effectiveness check look like?

- Draft the first paragraph of the Form 483 response to the CAPA observation. Include correction, corrective action and a systemic check.

### Table notes

### One thing I will change at my site

# Case 4: Three sites, three QMSs, one certificate

## Background

- Multinational with US headquarters (design and manufacturing), a German site acquired in 2023 (legacy MDD manufacturer) and a high-volume plant in Costa Rica; about 1,800 staff.

- One MDSAP certificate covers all three sites; MDR certificate; Health Canada licences.

- Integration of the German site's QMS was planned "after the MDR transition".

## The audit

| When | What happened |
| --- | --- |
| Surveillance audit, German site | The MDSAP auditor samples complaints received in the previous 12 months. One complaint from a Canadian hospital reports a serious deterioration in a patient's health associated with the device. It arrived at the German site via the distributor. |
| Same day | The complaint was assessed under the German legacy complaint procedure. The reportability decision took 14 days; the report reached Health Canada on day 22. The Canadian limit is 10 days. |
| Same day | The headquarters procedure would have flagged the complaint for a 48-hour reportability triage. Neither site knew the other's process in detail. |
| Day 2 | Three versions of the complaint form are in use across the company. German trend data has never reached global management review. |
| Day 2 | The previous audit raised inconsistent complaint intake as a grade 2 nonconformity. The CAPA was closed with "training". |

## Exhibits

EXHIBIT 4A German site complaint procedure QM-WA-08 (excerpt, translated) Reportability to authorities is assessed in the monthly complaint review meeting, using the criteria of MEDDEV 2.12/1 and the Medical Devices Act (MPG).

EXHIBIT 4B Previous MDSAP audit, nonconformity 3 (grade 2) and CAPA closure Finding: complaint intake forms differ between sites and do not consistently capture the country in which the event occurred. Corrective action: complaint handling training delivered to the German site team (attendance 18/18). Effectiveness: training records verified. Status: CLOSED.

EXHIBIT 4C Global complaint KPI dashboard (Q2) On-time reportability decisions (within 48 h): Headquarters (US): 97% Costa Rica: 95% Germany: n/a - not tracked

## Discuss at your table

- Why did "training" fail as a corrective action? What was the actual root cause?

- What must a single global complaint process contain to work in Germany, Costa Rica and the US at the same time?

- Exhibit 4A: identify every element of this clause that is incompatible with a multi-market QMS in 2026.

- Who should own the regulatory requirements matrix, and who should own keeping it current? What is the change-control link?

- Work out the MDSAP grade for the reporting nonconformity using the rules on page 1, and explain who hears about it and when.

### Table notes

### One thing I will change at my site

# Across the four cases

Look back at all four cases before the closing section of the seminar.

- Which failure appears, in some form, in every case?

- In each case, where was the earliest moment the outcome could have been changed, and who had the authority to change it?

- Which of the four companies would you least like to be, and why?

- Rank your own site's exposure to each of the four scenarios: low, medium, high.

| Case | Exposure at my site (low / medium / high) | First action |
| --- | --- | --- |
| 1. Unannounced notified body audit at the main site |  |  |
| 2. Unannounced audit at a critical supplier or subcontractor |  |  |
| 3. First FDA inspection under the QMSR |  |  |
| 4. Multi-site complaint handling and reporting |  |  |

### My list of things to change